## Slide 1
FINAL-YEAR DEMONSTRATION
Low-Cost
Myoelectric Hand Control
An accessible EMG-controlled prosthetic alternative
Team [Member 1] [Member 2] [Member 3] [Member 4]
UCL Biomedical Engineering · 2026

## Slide 2
01 · THE PROBLEM
Commercial Myoelectric Prosthetics
Are Out of Reach
2M+ upper-limb amputees worldwide
Commercial myoelectric hands: £30,000–£60,000
80% of amputees in low-income regions go without
Long NHS assessment & funding waiting lists
£30,000+
starting cost of a commercial
myoelectric prosthetic hand
Sources to add — Open Bionics, NHS prosthetics
2 / 8
EMG Robotic Hand · UCL BME

## Slide 3
02 · OUR SOLUTION
5-Gesture EMG Hand for ~£150
End-to-end signal chain: skin to servo
Forearm
electrodes
▶
AD8232
amplifier × 2
▶
ESP32
WiFi UDP
▶
Raspberry Pi 5
CNN inference
▶
PCA9685
servo driver
▶
5-finger
servo hand
Real-time (20 Hz)
Wireless wearable
5 distinct grasps
Same hardware reused as a live ECG / vital-signs dashboard — see Results slide.
3 / 8
EMG Robotic Hand · UCL BME

## Slide 4
03 · HARDWARE
Wearable + Edge Compute
Component | Qty | ~£
ESP32 DevKit V1 | 1 | 10
MyoWare 2.0 sensor + Cable Shield | 1 | 50
Raspberry Pi 5 (8 GB) | 1 | 70
PCA9685 PWM driver | 1 | 10
Micro servos (MG90S) | 5 | 20
5 V PSU (≥ 3 A) | 1 | 12
3× FSR + electrodes + accessories | — | 30
Total |  | ≈ £202
Photo of wearable + hand
[ insert image here ]
4 / 8
EMG Robotic Hand · UCL BME

## Slide 5
04 · SOFTWARE PIPELINE
Signal → Gesture → Action
Raw EMG
1 kHz · 1 ch · MyoWare 2.0 SIG
▶
RMS threshold
envelope · 5 s auto-calibration · debounced
▶
Decision
flex → close · relax → open
▶
Force feedback
3× FSR-402 halts closing per finger
▶
Servo angles
5 channels → finger positions
< 50 ms end-to-end latency on Pi 5 CPU · sample → prediction → servo command
5 / 8
EMG Robotic Hand · UCL BME

## Slide 6
05 · DEMO
Live Demonstration
▶
Embedded demo video — 45–60 s
Real-time gesture classification driving the servo hand
6 / 8
EMG Robotic Hand · UCL BME

## Slide 7
06 · RESULTS
Performance
rest
fist
open
pinch
point
predicted →
rest
fist
open
pinch
point
← true
99
0
1
0
0
1
95
1
3
0
2
2
93
1
2
0
5
0
78
17
0
0
3
22
75
Confusion matrix (placeholder — replace with your actual numbers)
92%
Per-subject
accuracy
80%
Cross-subject
accuracy
47 ms
End-to-end
latency
Baseline: 1D CNN vs SVM on classical features (MAV, RMS, ZC, WL) — CNN +8% accuracy on the same test set.
7 / 8
EMG Robotic Hand · UCL BME

## Slide 8
07 · IMPACT
£200 vs £30,000
150× cost reduction · open-source · replicable by BME students, makers, low-income clinics
FUTURE WORK
Multi-channel EMG
Distinct gestures beyond open/close
(currently 1 channel)
5 fingertip force sensors
Per-finger closed-loop control
(currently 3 FSRs)
Cross-subject pretraining
Generalise to new users
(Ninapro DB2 transfer)
Accessible myoelectric prosthetics, built today.